## Slide 1
Single Photon Detection with the Multi-anode CLAS12 RICH Detector
M. Contalbrigo – INFN Ferrara – on behalf of the CLAS12 RICH group
10th International Workshop on
Ring Imaging Cherenkov Detectors
31st July 2018,
Russian Academy of Science - Moscow
Notes: Vanno messi studi innovativi e pubblicazioni

## Slide 2
2
Contalbrigo M.
Cherenkov angle (mrad)
7 Gev/c
6 Gev/c
8 Gev/c
p
k
INSTITUTIONS
INFN (Italy) Bari, Ferrara, Genova, L.Frascati, Roma/ISS
Jefferson Lab (Newport News, USA)
Argonne National Lab (Argonne, USA)
Duquesne University (Pittsburgh, USA)
George Washington University (USA)
Glasgow University (Glasgow, UK)
Kyungpook National University, (Daegu, Korea)
University of Connecticut (Storrs, USA)
UTFSM (Valparaiso, Chile)
p
Goal kaon-pion separation up to 8 GeV/c (prototype results):
10th RICH Workshop, 31st July 2018, RAS - Moscow
CLAS12 RICH
Notes: Vanno messi studi innovativi e pubblicazioni

## Slide 3
3
Photon Sensor: MA-PMT
MA-PMT
< 1 cm spatial resolution
< 1 ns time resolution
Compatible with the low torus fringe field
Average MA-PMT gain ~ 2.7 106
Corresponds to SPE ~ 400 fC
✓ 64 6x6 mm2 pixels cost effective device
✓ High sensitivity on VIS towards UV light
✓ Mature and reliable technology
✓ Large Area (5x5 cm2)
✓ High packing density (89 %)
✓ Fast response
✓ Expensive technology
10th RICH Workshop, 31st July 2018, RAS - Moscow
~ 1 m2 active surface
Notes: Dire meglio chi fa cosa.

## Slide 4
4
RICH Readout Electronics
Contalbrigo M.
Readout Electronics
Compact (matches sensor area)
Modular Front-End (Mechanical adapter, ASIC, FPGA)
Scalable fiber optic DAQ (TCP/IP or SSP)
Tessellated (common HV, LV and optical fiber)
10th RICH Workshop, 31st July 2018, RAS - Moscow
HV
LV
Optical
Fiber
SSP Fiber-Optic DAQ
Tile power dissipation ~ 3.5 W
Notes: Immagine electronic panel e scheda ottica

## Slide 5
5
Contalbrigo M.
Analog: Charge (1 fC)
Digital: Time (1 ns)
Trigger latency (8 ms)
Optical ethernet (2.5 Gbps)
Trigger: external
internal
self
On-board pulser
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Charge
Test
Test pulse
Slow shaper
Slow shaper
Digital pulse
Linear response
Multiplexed readout
Limited holding time delays
Used for calibrations
RICH Front-End Electronics
Notes: Aggiungere range di validita’ per fattorizzazione GPDs

## Slide 6
6
ADC Charge Measurement
Multiplexed readout up to 50 kHz
High resolution SPE spectrum
Viable for efficiency and gain monitors
In conjunction with timing, allows the
study of PMT discharge and cross-talk
Contalbrigo M.
Gain
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Dire meglio chi fa cosa.

## Slide 7
7
Contalbrigo M.
Analog: Charge (1 fC)
Digital: Time (1 ns)
Trigger latency (8 ms)
Optical ethernet (2.5 Gbps)
Trigger: external
internal
self
On-board pulser
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Start time
Stop time
Test
Test pulse
Slow shaper
Slow shaper
Digital pulse
Digital response
Working in saturated regime
64 parallel channel readout
8 ms FIFO and delays
1 ns time resolution
RICH Front-End Electronics
Notes: Aggiungere range di validita’ per fattorizzazione GPDs

## Slide 8
8
TDC Digital Readout
Pedestal rms as seen by a test-point
As seen by RICH readout
During Acceptance tests
During Internal Pulser Calibration
Discrimination down to 20 fC, i.e. few % of SPE, allows sensor characterization
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Vedere presentazione Matteo

## Slide 9
9
Sensor
Electronic
Gain x 4
THR ~ 1/80 SPE
Signal
Cross-talk
Sampling time
10th RICH Workshop, 31st July 2018, RAS - Moscow
Optical and Electronic Cross-talk
Notes: Dire meglio chi fa cosa.

## Slide 10
Threshold
Eff. ratio
0 0.5 1 2 4
Eff. ratio
Maroc
Gain
Single-photon Discrimination
Relative efficiency map
Relative gain map
10
Contalbrigo M.
x ¼
x ½ MAROC
x 1 GAIN
x 2
x 4
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Mettere relazione con blue luminosity ?

## Slide 11
11
Single Photon Electron Timing
± 3ns
Typical time-walk with charge
Channel by channel time calibration: no offsets no walk
Laser
PMT
Time (ns)
Occupancy (#)
Contalbrigo M.
Time over threshold relates to charge
Duration [ns]
Tme [ns]
ADC amplitude [DAC]
Time over THR [ns]
1
Cherenkov
angle
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 12
12
RICH Installation
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 13
13
Electronic Pedestals
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Pedestal rms [DAC unit]: without and with grounding grid
Jan 18
-- off
-- 1000 V
Mar 18
-- off
-- 1000 V
Pixel #
Pixel
Pixel #
Pixel #
Pixel
Pixel
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 14
14
Online Equalization
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
After equalization: distributions narrower and less sensitive to the common threshold
saturate signals and cross-talk well separate
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 15
15
Single Photon Time Analysis
Contalbrigo M.
Dark counts
Other bunch
10th RICH Workshop, 31st July 2018, RAS - Moscow
CLAS12 Reconstructed Time and Position:
Photons are traced using information
from other CLAS12 detectors
RICH Measured Time and Position:
Defined by the RICH DAQ
Good photons should match
in time and space
Time analysis allows to
separate spurious signals
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 16
16
Time Calibration
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Time vs Duration
after offset correction
Time vs Duration
after walk correction
Cross-talk
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 17
17
Single-photon Time Resolution
Contalbrigo M.
PMT 270
s = 0.58 ns
PMT 52
s = 0.66 ns
10th RICH Workshop, 31st July 2018, RAS - Moscow
before time-walk correction
after time-walk correction
Single-photon time resolution better than the 1 ns specification
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 18
18
Cherenkov Photon Reconstruction
Contalbrigo M.
Example of signal hits identified by time consistent with CLAS12 reconstruction
10th RICH Workshop, 31st July 2018, RAS - Moscow
Direct ring
Half recflected ring
Next
Page
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 19
19
Direct Cherenkov Photons
Contalbrigo M.
About 18 photons for a center
ring (no reflection accounted for)
Consistent with the TDR projection
Preliminary single photon Cherenkov
angle resolution = 6 mrad
Close to the 4.5 mrad goal despite
No alignment
Nominal (no real) optical property
10th RICH Workshop, 31st July 2018, RAS - Moscow
In principle enough for a Cherenkov resolution at track level of s = s1/√N ~ 1.5 mrad
s = 8 mrad
s = 6 mrad
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 20
20
Application: Modular RICH @EIC
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Aerogel
Fresnel lens
Sensor plane
Compact and modular RICH indipendent elements
H13700 to reach the 3 mm spatial resolution
See Xiaochun He talk
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 21
21
Application: SiPM Arrays
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Pixel #
Pixel #
Peltier cell cooling
Down to -30o in N2
Notes: Fare una immagine del PMT sulla risoluzione

## Slide 22
22
Contalbrigo M.
Application: DIRC @ GlueX
H12700 +
CLAS12
readout
Hadron discrimination up to 4 GeV/c
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Dire meglio chi fa cosa.

## Slide 23
CLAS12 RICH designed to provide hadron identification in the 3 to 8 GeV/c momentum range
A hybrid-optic design has been adopted to minimize the instrumented area to about 1 m2
Flat-panel multi-anode PMTs are being used for the first module
SiPMs are being investigated for the second module
The readout electronics is designed to offer
Discrimination down to few % of SPE
Time resolution of 1 ns
Negligible dead time at 30 KHz
Trigger latency up to 8 ms
Featuring:
Compatibility with various sensors and applications
Modular Front-End (Mechanical adapter, ASIC, FPGA)
Scalable fiber optic DAQ (TCP/IP or SSP)
Compact and tessellated geometry (common HV, LV and optical fiber)
Flexible trigger logic (external, auto, self)
Charge measurement (multiplexed ADC or time-over-threshold)
Multi purpose electronics: in use also for GlueX DIRC and EIC R&D
23
Conclusions
Contalbrigo M.
10th RICH Workshop, 31st July 2018, RAS - Moscow
Notes: Dire meglio chi fa cosa.